# General Camera Instructions

## Setting up the camera

- Set up the tripod at a reasonable height to image the patient. Make sure that the bubble on the level is in the center.

- Set camera up on tripod by attaching it to the base plate using the screw to secure it so it does not move.

- Turn on and log into laptop

- Plug one end of the USB into the laptop and the other end into the camera

- Open the IR Flash software on the laptop

- *You should be able to see a live image showing up

- Note: Once IR Flash is open DO NOT unplug the camera until you close the software!

- Make sure the Palette is set Gray Scale

- Make sure the AGC box us unselected

- Adjust the Min and Max values based on the histogram to have the best contrast

- *Generally, min around 23 and max around 38

- Focus the camera by moving the camera so that the distance between the patient and the front of the lens is between 14-21 inches

* Generally around 17 inches is a good focus

* Adjust based on the patient

* Only grab the camera body if you absolutely have to. After touching the camera body give it a minute or two to get stabilized again

10. After focusing, set the Averaging to 30

Note: If the live image ever starts to randomly get dark or have a weird gradient, go to Camera  Image Touchup. This would likely happen while the camera is warming up. Give the camera at least 15 minutes to warm up.

## Taking an Image Series

- Select Video  Configure Series Image Capture

- Set Capture Type to “Capture for fixed time”

- Set Time to 0 hours 15 minutes 0 seconds.

- Change Capture Rate to 2 Per Minute. Select OK

- Select Video  Capture Image Series. Under the live image

- feed you should now see a red record button.

- To begin the image series click the red record button:

- To save all the images from the series, select

- File  Save All and make sure to save as a Raw 16 Bits

TIF File (Note: This will save all the images that are shown in the Image bar with numbers starting from NAME-0001 regardless of the capture number).

## Taking a Single Image

- Select Video  Capture Single Images

- To take an image click the camera button. It will be added to the Image bar shown below.

- To save the image, select the image and click Save Image